LINEFIT ver. 14.8 - documentation of upgrades over ver. 14.5

Date: 2021-Apr-21, Version 1.0

The last version of the code distributed among NDACC-IRWG and TCCON operators was version 14.5. In version 14.6, the previously used speed-dependent Galatry lineshape model was replaced by the Hartmann-Tran lineshape model and first (Rosenkranz) and second order (Smith) line mixing capabilities have been added. This version was not generally distributed but used for lab spectroscopy applications by DLR (M. Birk, G. Wagner) and KIT (D. Dubravica).

A test version with new capabilities requested by TCCON was handled as ver. 14.7, this current LINEFIT ver. 14.8 is the consolidated version of ver. 14.7.

Acknowledgements: Special thanks to Manfred Birk and Georg Wagner for using the software, carefully testing, and highly useful feedback! When using LINEFIT on ozone spectra, Manfred and Georg discovered a bug in the “species.inf“ file for the ozone isotopologues, which contained wrong masses – this error is corrected from version 14.6 onwards. Special thanks also to Sebastien Roche (Univ. of Toronto), Mahesh Kumar Sha and Bavo Langerok (both BIRA) for exploring ways to make LINEFIT output visually user-accessible by providing tools for visualisation and batching. The ALIGN visualisation tool provided by Dan Smale and colleagues (NIWA) is again included in the new version 14.8 distribution. I would like to thank Roland Ruhnke and Lena Feld (both KIT) for support in creating iOS and Linux versions of the code and for extensive testing of the latter. Finally, I would like to thank David Griffith (Univ. Wollongong) for continuously pushing towards a more user-friendly behaviour of the code for TCCON. This update is essentially in response to his request.

TCCON prescribed mode for HCl cell analysis: The new prescribed TCCON mode requires from the user only (1) provision of a spectrum recorded with OPDmax = 45 cm or higher, of sufficient spectral coverage (5670 – 5805 cm-1) and sufficient SNR (~ 2000) and (2) to apply the demo input file located in folder examples\TCCON\lamp\HCl\prescribed. In this file, only the name of the spectrum file, the internal semi-field-of-view (apt / (2 * 418 mm)), and the correct temperature T and pressure values (note: p_tot and p_self should be equal) have to be specified by the user. The reference p values for each cell are provided in the table of the TCCON wiki (note: a change of T will also affect the effective pressure values. If the temperature deviates from the reference T = 296 K of the table, the pressure values need to be updated according to the gas law). The table can be found here: https://tccon-wiki.caltech.edu/index.php?title=Software/LINEFIT/Effective_pressure_values_for_HCl_cells&highlight=hcl+cells

When the code is started, the compatibility of input file settings with the prescribed mode is tested, if a deviation is found, the execution is stopped with a specific message.

Next, LINEFIT investigates whether the spectral coverage is sufficient and resamples the input spectrum. In this step, the resolution is limited to 45 cm. The chosen sampling width (factor 8 oversampling) allows a visual inspection of retrieval results. A common choice for the spectral sampling of the fit is also useful as this interacts with the interpretation of constraints (as a larger number of spectral gridpoints will increase the spectral contribution to the cost function).

The SNR is tested on the resampled spectrum (range used: 5715.0 - 5716.5 cm-1). The SNR deviation is determined wrt a local reference value, which is determined by applying a low-pass filter on the spectrum for creating the local reference (convolution with a window function of width ~0.14 cm-1). This SNR is slightly too optimistic, as it suppresses the spectrally broader contributions (extrapolation towards a total SNR based on a white noise assumption is not performed). The high-pass SNR has been chosen as a measure of spectra quality because it allows spectra suffering from not too-high frequency channeling to pass th quality check. However, at the end of the inversion, the total SNR calculated from the fit residual is compared with the SNR derived from the measurement as described above. An error message is displayed if the fit SNR is below 70% of the previously determined high-pass SNR. In this case the user needs to check  the measured spectrum, probably it contains channeling, which should be included when repeating the fit (for this purpose, the number of resonators to take into account and their channling frequencies need to be determined from the fit residuals by, e.g., repeating the fit selecting ILS_choice = 3 (extended) instead of TCCON mode).

In the next step, the mean value in the (5715.0 - 5716.5 cm-1) window is used for normalizing the measured spectrum to unity. A global scaling factor on the measured spectrum is applied - slope, curvature, etc. are imposed on the calculated spectrum during the retrieval (baseline parameters). The normalisation of the background level is useful as this value interacts with the interpretation of constraints (higher values will increase the spectral contribution to the cost function).

Before the retrieval is started, the spectral calibration of the measured spectrum is checked and modified, if required. A classical least-squares fit algorithm tends to fail when measured and calculated spectral lines do not overlap any more due to a miscalibration of the measurement. For this reason, LINEFIT checks the spectral calibration of the measurement before the retrieval is started. In normal operation mode, the execution stops when the spectral detuning found is too big. This is because a general spectral scene might contain a dense forest of spectral lines, and allowing for larger range of the recalibration might lead to a misassignment of lines. Therefore, the problem of checking the proper calibration of the measured spectrum is forwarded to the operator. In TCCON prescribed mode, the spectral scene is well-determined and simple. The HCl line located at 5739.262511 cm-1 is used for a pre-adjustment of the spectral calibration parameters before the retrieval is started. The line intensity is also evaluated. If the line is significantly weaker than expected for a TCCON cell, the reported flag status will differ from zero.

The global constraints (smoothness of modulation amplitude and phase) to be used by the retrieval are adjusted by the code using the relation reg = 0.1 * (9000.0d0 / (hcl_amp * high_pass_SNR))**2, here hcl_amp is the center absorption strength of the HCl line located at 5739.262511 cm-1 (the absorption strength of a resolved, fully opaque line would be 1.0). Note that the phase error is clamped to zero at ZPD (zero path difference) for the TCCON prescribed mode. This is equivalent to the assuming the phase correction scheme used for the Fourier transform of the interferogram works perfectly. The modulation retrieved from the cell measurement will likely be incorporated in the GGG processing scheme in some future. As the atmospheric lines of interest are much broader than the HCl lines, their retrieval depends mainly of the modulation efficiency assumed for the section of the interferogram around ZPD.

If the SNR determined from the fit residual is worse than expected, the reported flag status will differ from zero. In TCCON prescribed mode, an additional output file TCCON.dat is generated, which among other items reports the flag status, the high pass SNR and the global constraints used for the retrieval.

Commented LINEFIT input file, required user-specific settings highlighted in red.

Input file Linefit ver. 14.0

input

species files (masses and partition sums)

number of species

for each species: linelist (HITRAN format)

$

hit\species.inf

2

hit\15_hit08.par

1

151

hit\15_hit08.par

1

152

output

Output directory

Working path

Output grade (0: normal, 1-3: extended)

$

ergs

wrk

0

target of program call: select from...

1:forward calculation for specified spectral regimes

2:forward calculation for given measured spectra

3:ILS retrieval

4:Fit of aux parms only (fixed ILS parameters)

$

3

ILS model:

1: Use simple ILS model

2: Use TCCON ILS model

3: Use extended ILS model

4: TCCON HCl cell analysis (extended IlS + prescribed parameter settings)

$

4

Fit only real spectrum (false) or also fit imag spectrum (true)?

(the fit of an imag spectrum requires extended ILS model)

Assume single-sided (false) or double-sided IFG (true)?

(the double-sided option requires fit of imag spectrum true)

Assume common spectral abscissa for all species (T/F)?

$

.false.

.false.

.false.

number of microwindows (max. 26)

$

1

1.0d-3

(5675.00,5800.20)

deweighting regions for each MW

$

16

(5677.8,5679.1),0.001

(5689.0,5692.6),0.001

(5696.9,5698.1),0.001

(5708.6,5709.9),0.001

(5721.9,5725.8),0.001

(5732.0,5733.0),0.001

(5740.8,5743.6),0.001

(5747.3,5748.1),0.001

(5757.0,5758.4),0.001

(5759.5,5762.1),0.001

(5771.5,5774.2),0.001

(5783.5,5785.6),0.001

(5792.6,5793.4),0.001

(5797.1,5797.9),0.001

(5798.3,5798.8),0.001

(5799.2,5799.6),0.001

Resampling of measured spectra? (T/F)

Sampling density (integer value: 1 minimally sampled, >1 oversampling)

(Note: Resampling will limit the resolution of the measured spectrum

according to the OPDmax value specified below)

$

.true.

8

(in case 2/3)

format of spectra (0: dpt, 1: bin, 2: OPUS transmittance)

measurement files (one for each microwindow)

(if fit of imag spectrum:

give all real spectra, then all imag spectra)

$

2

spectra\hcl-tccon\081009-bremen\hcl081009-vac.0

species parameters:

for each species:

gas T,fit of gas T (F/T),column of species [m-2], ptot [mbar], fit of total pressure (F/T), ppart[mbar], default gamma

(cell column = 7.243e24 * p[mbar] * l[m] / T[K])

(first-guess values in case of retrieval)

for each MW: take species into account(T/F),column scaling factor, spectral scaling factor of species - 1

$

294.0,.false.,1.268e22,5.10,.false.,5.10,0.0075

.true.,1.0,0.0

294.0,.false.,1.268e22,5.10,.false.,5.10,0.0075

.true.,1.0,0.0

MW dependent parameters I: (first-guess values in case of retrieval)

for each microwindow:

number of baseline parameters, values of baseline parameters

$

5  1.0  1.0  1.0  1.0  1.0

MW dependent parameters II: (first-guess values in case of retrieval)

for each microwindow and species

number of spectral abscissa parameters, values of spectral abscissa parameters ...

$

MW1

8  0.0  0.0  0.0  0.0  0.0  0.0  0.0  0.0

8  0.0  0.0  0.0  0.0  0.0  0.0  0.0  0.0

channeling (first-guess values in case of retrieval)

number of channeling frequencies to be handled (same for all mws)?

frequencies: period [cm-1] (same for all mws!)?

$

0

channeling amplitudes

for each mw: (amplitude cosine,amplitude sine)

frequency--->

m

w

$

(0.0d0,0.0d0)

(if forward calculation ... case 1/2)

Should noise be added to calculated spectra?

standard deviation of instrumental noise in unapodised spectrum

$

.false.

0.0

instrumental parameters:

apodisation

(1: boxcar

2: triag

3: Hamming

4/5: Blackmann-Harris 3-term/4-term

6/7/8: Norton-Beer weak/medium/strong)

maximal optical path difference [cm]

Should self-apodisation taken into account?

maximum inclination of rays in the interferometer [rad]

(aperture radius / focal length of collimator)

$

1

45.00

.true.

1.196e-3

gas cell parameters:

handle self emission:

0     pure gas transmission (no self emission)

1/-1  transmission cell spectrum with self-emission (input / output port cell)

2     calibrated emission spectrum (nW / cm2 sterad cm-1)

3     laser lines

cell window temperature [K]

black-body temperature [K]

black-body emissivity

external transmittance of each cell window

$

0

294.0

1273.0

0.98

0.8

adjustment of internal spectral grid:

adjustment of internal spectral grid width (default: 1.0)

mw-extension for calculation of irradiated transmission spectra

(equivalent to width of ILS folding kernel, pos. integer, in multiples of 1/2*opdmax)

maximal allowed spectral shift (pos. integer, in multiples of 1/2*opdmax)

$

1.0

100

20

number of iterations in case of retrieval

convergence threshold (positive, value nearer to zero triggers more iterations)

L-M stab value

$

10

1.0d0

0.0d0

Reg Modulation (dev+smooth), Reg Phase (dev+smooth),reg T(dev)

Reg column (dev+smooth, per species), reg spectral scale (mean, per species)

$

1.0e4,0.3,1.0e4,0.3,0.0

0.0,0.0,0.0001

0.0,0.0,0.0001

Parameter apo, Parameter Phase

(if ILS model simple, then only first pair is used

and interpreted as linear modulation and constant phase error)

(if ILS model extended: 20 / 39 values)

$

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

Modulation profile used for regularisation

(only used in case of extended ILS model: 20 / 39 values)

$

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

1.0,0.0

Modulation: Position-dependent regularisation strength (dev+smooth)

(only used in case of extended ILS model: 20 values)

$

1.0000,1.0000

0.0000,0.9955

0.0000,0.9822

0.0000,0.9603

0.0000,0.9305

0.0000,0.8936

0.0000,0.8504

0.0000,0.8021

0.0000,0.7498

0.0000,0.6945

0.0000,0.6376

0.0000,0.5801

0.0000,0.5231

0.0000,0.4674

0.0000,0.4140

0.0000,0.3633

0.0000,0.3160

0.0000,0.2724

0.0000,0.2327

0.0000,0.1970

Phase: Position-dependent regularisation strength (dev+smooth)

(only used in case of extended ILS model: 20 values)

$

1.0000,1.0000

0.0000,0.9955

0.0000,0.9822

0.0000,0.9603

0.0000,0.9305

0.0000,0.8936

0.0000,0.8504

0.0000,0.8021

0.0000,0.7498

0.0000,0.6945

0.0000,0.6376

0.0000,0.5801

0.0000,0.5231

0.0000,0.4674

0.0000,0.4140

0.0000,0.3633

0.0000,0.3160

0.0000,0.2724

0.0000,0.2327

0.0000,0.1970

Contents of TCCON.dat file:

Flag status:   0

Possible flagging results:

0: fit conforms to TCCON requirements

1: fit residuals higher than expected

10: HCl column amount out of range

100: measured spectrum has insufficient SNR

HCl line (5739.26 cm-1) amplitude:  5.676E-01

Preadjusted spectral detuning:  3.574E-01

SNR in measured spectrum:  8.599E+03

SNR achieved in fit:  8.008E+03

REG_APO2 REG_PHAS2:  3.399E-01 3.399E-01