## Slide 1
Detection of CO2 emitted by the megacity Berlin using an array of portable FTIR spectrometers

## Slide 2: Outline
COCCON – network
Instrumentation
Berlin campaign
Summary and outlook

## Slide 3: COCCON – network
Collaborative Carbon Column Observing Network
Initiated by KIT
  EM27/SUN developed by KIT (M. Gisi, F. Hase et al,) in cooperation with Bruker Optics, Ettlingen as starting point
  Superb instrumental performance with the prototype spectrometer – results comparable to TCCON reference
 Novel compact NIR-FTIR spectrometer sold by Bruker as EM27/SUN
COCCON aim:
  Use the mobility of EM27/SUN as a supplement of the TCCON
  Quantification of the sources and sinks of CO2 and CH4
  Increase the density of ground-based XCO2 and XCH4 observations
  Satellite validation, input data for global Carbon cycle models
COCCON start – implement 10 EM27/SUN spectrometers
  Recalibration and service in Karlsruhe (availability of TCCON spectrometer and Bruker support)
  Spectrometer operation in cooperation with local groups
  Emphasis on local capacity building

## Slide 4: COCCON: Instrument overview
EM27/SUN spectrometer from Bruker
RockSolidTM pendulum interferometer
MOPD: 0.9 cm; Resolution: 0.5 cm-1, double sided
InGaAs (Indium Gallium Arsenide) detector used
Spectral range: 5000 – 11000 cm-1
Field-of-view: 2.36 mrad
Standard non freq.-stabilized HeNe reference laser
Dimensions: 47 x 63 x 35 cm
Mass: ~ 25 kg including tracker
Sun tracking using optical
feedback by the camtracker
software
EM27SUN spectrometers measuring at IMK
Fig. Source: Bruker

## Slide 5: Performance of the EM27/SUN
Scaling retrieval of O2, CO2, CH4 and H2O using PROFFIT
Long term stability of the EM27/SUN XCO2 < 0.1%
Side-by-side intercomparison
Calibration factors before / after campaign stable within 0.005 % for XCO2 and 0.025% for XCH4  see talk from Matthias Frey
Calibration with respect to TCCON reference
EM27/SUN spectrometers measuring at IMK

## Slide 6: Encircling source area of interest – megacity Berlin
Why Berlin? – German capital, well-defined source, flat terrain
Campaign timeline: 23.06.2014 – 11.07.2014 (3 weeks)
Participants:
  Thomas Blumenstock, Matthias Frey, Jochen Groß, Frank Hase, Matthäus Kiel, Mahesh Kumar Sha, Gizaw Mengistu Tsidu
  Support: Regina Kohlhepp (DWD), Klaus Schäfer (IMK-IFU)
Fig source: aveato.de
Fig source: dw.de

## Slide 7: Berlin campaign: sites
5 EM27/SUN spectrometers distributed in the periphery around Berlin
Site | Long. (° E) | Lat. (° N) | Alt. (m a. s. l.)
Mahlsdorf | 13.589 | 52.486 | 39.0
Charlottenburg | 13.302 | 52.505 | 47.7
Heiligensee | 13.228 | 52.622 | 34.5
Lindenberg | 13.519 | 52.601 | 63.3
Lichtenrade | 13.392 | 52.391 | 44.8

## Slide 8: Berlin campaign
Calibration of spectrometers before / after the campaign
Careful site selection – free line-of-sight towards sun
Auxiliary meteorological information
  Local pressure, temperature
  Ceilometer measurements: BLH as a function of time*
  COSMO winds**
*: Prof. K. Schäfer, KIT, IMK-IFU
**: R. Kohlhepp, DWD

## Slide 9: Berlin campaign
XCO2 time series
F. Hase et al., “greenhouse gas emissions from Berlin – part 2 ...“, AMTD, 2015

## Slide 10: Berlin campaign
Simplified source strength of Berlin (total source strength: 730 kg CO2 / sec)
F. Hase et al., “greenhouse gas emissions from Berlin – part 2 ...“, AMTD, 2015

## Slide 11: Berlin campaign
Comparison with simplified model – June 27, 2014
F. Hase et al., “greenhouse gas emissions from Berlin – part 2 ...“, AMTD, 2015

## Slide 12: Paris campaign
5 EM27/SUN spectrometers IMK – KIT: 3 EM27/SUN sites LSCE – IPSL: 1 EM27/SUN site LERMA – UPMC:1 EM27/SUN site + Bruker IFS 125HR LISA – UPEC: mobile support
Side-by-side intercomparison at KIT

## Slide 13: Paris campaign
XCO2 time series
07.05.2015
10.05.2015
11.05.2015
Preliminary
Calculations: M. Frey

## Slide 14: Summary and outlook
Superb performance of the EM27/SUN as a mobile FTIR spectrometer measuring GHGs
Successful city campaign estimating the source concentration of CO2 from the megacity Berlin
City campaign performed in Paris in collaboration with local groups
Add-on to the TCCON network
Differential observation of gradients, local sources
Observations from moving platforms
COCCON sites | Personel
Gobabeb | Ruusa Gottlieb
Addis Ababa | Gizaw Mengistu
Gadanki | Alok Taori
Volta Redonda | Carlos Fellows (?)
Los Alamos | Manvendra Dubey
Harvard | Jai Chen
... | ...

## Slide 15
If you are interested to participate in this effort, feel yourself invited.
Please contact:
Dr. Frank Hase, KIT (frank.hase@kit.edu) – COCCON PI
Dr. Thomas Blumenstock, KIT (thomas.blumenstock@kit.edu)
Dr. Mahesh Kumar Sha, KIT (mahesh.sha@kit.edu)
Matthias Frey, KIT (m.frey@kit.edu)
Thank you – Danke schön – शुक्रिया

## Slide 16: Extra sildes

## Slide 17: Solar spectrum measured by EM27/SUN

## Slide 18: Spectral window used by PROFFIT
Examples of
Spectral fit
Spectral range used for
retrieval of gases:
  O2: 7765.0 – 8005.0 cm-1
  H2O: 8353.4 – 8463.1 cm-1
  CO2: 6173.0 – 6390.0 cm-1
  CH4: 5897.0 – 6145.0 cm-1

## Slide 19: COCCON
Results achieved with the prototype were exciting
M. Gisi, F. Hase, S. Dohe, T. Blumenstock, A. Simon, A. Keens: “XCO2 measurements with a tabletop FTS …“ AMT, 2012

## Slide 20: O2 column versus in-situ pressure time series
Data scaled to 30 m a. s. l.
In-situ factor ~ 3 %
  Spectroscopic uncertainties
  Commonly observed
Column data fits in-situ observations
Courtesy: M. Frey