## Slide 1
TCCON extension to lower wavenumbers - InGaAs-InSb setup and first mid-infrared retrievals from the sites Bremen, Nicosia and Orleans
Christof Petri, Constantina Rousogenous, Winfried Markert, Thorsten Warneke and Justus Notholt

## Slide 2
Plan: extend the measurements at our TCCON sites to lower wavenumbers -> simultaneous InGaAs-InSb measurements
Harmonize the setup and the retrieval strategie
Bremen and Nicosia are already set up, Orleans will follow but additionally needs automated LN2 filling.
TCCON extension to lower wavenumbers
Bremen Nicosia Orleans

## Slide 3
Si-detector
- not used anymore, but still ready on socket 4
InSb-detector
- installed on socket 1
Connection cable for dual channel recordings
on socket 1 and 2
InGaAs-detector
-on socket 2
5-plc. Filter wheel
- installed
50-50 beamsplitter
- in self constructed holder
Holder for HCl cell
Empty socket 3
8-plc. filter wheel
- placed in the interferometer chamber

## Slide 4
Instrument status before alignment:
- just inside TCCON requirements
Instrument status after alignment:
- almost perfekt
Centered fringes as seen with the new alignment tools

## Slide 5
3D printed parts:
  Beamsplitter-holder for the 2“ circular 50:50 beamsplitter
    Adjustable with fine tread screws
  Filter holder
    with handle
    wings, replacing the plastic washers
    different apertures to adjust the intensity of the different filters
3D files will be provided
TCCON extension to lower wavenumbers

## Slide 6
3D printed alignment tools:
  Camera tower with adjustable mirror and focus
  Holder for the alignment laser, adjustable with white screen to see the back reflection
Alignment tools

## Slide 7
Filter ranges
TCCON extension to lower wavenumbers
filter 7
filter 6
filter 5
long pass
4150 3500 3000 2500 1900

## Slide 8
Measurement schedule:
Time | Name | InGaAs | InSb filter | InSb-wavenumber | resolution
30min. | TCCON | open | Long pass | 1900 - 4150 | 0.014 cm-1
30min. | scd-2 | open | NDACC 7 | 1900 - 2230 | 0.014 cm-1
30min. | scd-3 | open | NDACC 6 | 1950 - 2650 | 0.014 cm-1
30min. | scd-4 | open | NDACC 5 | 2400 - 3200 | 0.014 cm-1
30min. | TCCON-LR | open | Long pass | 1900 - 4150 | 0.2 cm-1
3 x weekly | NDACC-set | ------ | 2, 3, 5, 6, 7 | 1900 - 4150 | 0.005 cm-1
TCCON extension to lower wavenumbers

## Slide 9
Retrieval schedule:
Name | spectra | InGaAs wn | InSb-wn | gases
TCCON | InGaAs only | 4000-11000 | ------- | TCCON
Long-pass | InSb + InGaAs | O2 only | 1900 - 4150 | C2H6, C2H2, CO, CO2, CH4, H2CO, HCN, N2O, NO
scd-2 | InSb + InGaAs | O2 only | 1900 - 2230 | CO, NO, O3
scd-3 | InSb + InGaAs | O2 only | 1950 - 2650 | CH4, N2O, OCS
scd-4 | InSb + InGaAs | O2 only | 2400 - 3200 | C2H6, C2H2, CO2, CH4, H2CO, HCN
TCCON-LR | InSb + InGaAs | tbd | tbd | 
NDACC-set | InSb only | ------- | 1900 - 4150 | NDACC
TCCON extension to lower wavenumbers

## Slide 10
First results: XC2H6 (2976-3305)
TCCON extension to lower wavenumbers
5

## Slide 11
TCCON extension to lower wavenumbers
First results: XN2O (2443-2813)
5

## Slide 12
First results: XCH4 (5938-6076 vs. 2599-2921)
TCCON extension to lower wavenumbers

## Slide 13
FTIR issues Nicosia
We found ‚longer than usual‘ scan durations ‚DUR‘ at Nicosia
At Orleans all ! measurements have precisely the same durations
This started already in 2019: 3% of measurements
In 2022: 27% !
-> filter out longer scan durations
212
211.5
211
212
211.5
211
2019: 3%
2022: 27%

## Slide 14
FTIR issues Nicosia
Scatter of MVD parameter increases in 2022
Filtering for the ‚too long‘ measurements: mvd looks much better in 2022.
-> additionally filter mvd > 0.75
Tests to measure with 20kHz, 10kHz, 5kHz

## Slide 15
QC-monitor at our stairway
  automatically updated
  using the touchscreen you can check the fast retrieval results from different sites.
QC-monitor

## Slide 16
Extended TCCON measurements are performed at our sites Bremen and Nicosia, Orleans will follow
First results are shown
MIR measurements:
  Fits have to be checked for tilt and curvature
  Correction factors for have to be found
Persistent instrumental issues at Nicosia
Summery and outlook

## Slide 17
Thank you
for
your attention!